## Slide 1
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
1. General principles and architecture
Global architecture
- CHARA Cosmic Debris is the highest client (including SPICA-FT server)
- OB supervisor for observation information (TBC -> Science Survey Management presentation)
- OS observation information from CD
- OS control of all SPICA-VIS devices thru servers

## Slide 2
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
1. General principles and architecture (continued)
SPICA-VIS Instrument Control Software (ICS)
- ICS is implemented as CHARA client/server applications (one per group of devices)
- Each server is registered to CHARA sockman
- An engineering GUI is available for each application
- Power control of each group of devices with remote controlled switches (power server)
SPICA-VIS Observing Software (OS) :
- Registers on CD messages
- Must be able to run on local control network or on Remote Observing Facility in Atlanta
Use of standard CHARA library and general encoding principles :
- Based on C language and GTK toolkit
- Use of CHARA messages architecture

## Slide 3
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
2. ICS: Active devices – Periscope and injection table

## Slide 4
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
2. Active devices – Periscope and injection table (continued)
+
+
Custom rotation stage
Lin engineering R256
Custom tip-tilt stage
with Picomotors
Newport 8740

## Slide 5
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
2. Active devices – Periscope and injection table (continued)
Custom rotation stage
Lin engineering R256
+
+
Newport 8740
Newport 8814 (TBC)
PI piezo TT
DNA cube
+
TBC

## Slide 6
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
2. Active devices – Periscope and injection table (continued)
+
Newport SMC100PP
Newport MFA-PP
Thorlabs MFF101/M

## Slide 7
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
2. Active devices: summary
Modules | Motors | Number | Type | Controller | Server
SFO (SPICA Feeding Optics) | FOP | 1 | DC motor | VEGA controller | spica_periscope_server
SFO (SPICA Feeding Optics) | TT-n / TP-n n∈[1,6] | 12 | Custom TT stage with Picomotors | Newport 8742 | spica_ip-ctrl_server
PDC (Polarization Delay Compensator) | PDC-n n∈[2,6] | 5 | Rotation Stage with Stepper Motor | R256 – Lin Engineering | spica_pdc_server
ADC (Atmospheric Dispersion Compensator) | AF-n / AS-n n∈[1,6] | 12 | Rotation Stage with Stepper Motor | R256 – Lin Engineering | spica_adc_server
BAS (Beam Adjustment System) | PH-n / PV-n n∈[1,6] | 12 | Newport 8814 (TBC) | Newport 8742 | spica_ ip-ctrl _server
BAS (Beam Adjustment System) | FT-n / FP-n n∈[1,6] | 6 | PI | DNA Cube | spica_tt-tracker_server
SFI (Spatial Filter Input mode) | DL-n n∈[2,6] | 5 | Newport MFA-PP | Newport SMC100PP | spica_delayline_server
BSL (Beam Selector) | SH-n n∈[1,6] | 6 | Thorlabs MFF101/M |  | spica_shutter_server
BCS (Beam Control System) | BCC | 1 | Newport MFA-PP | Newport SMC100PP | spica_ipdet_server
DIS (Dispersion System) | grating wavelength | 2 (3) TBD | TBD | TBD | spica_grating_server
FBI (Fiber Back Illumination) | FBI | 1 | TBD | TBD | spica_power_server
Total number of motors:
Injection table: 60
Spectrograph: 3 (or 4)
Other devices:
- 2 cameras
- Retrofeed source
- Spectral source(s)

## Slide 8
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
3. ICS: Low-level servers and GUIs
Servers
spica_periscope_server
spica_adc_server
spica_pdc_server
spica_ip-ctrl_server
spica_tt-tracker_server
spica_delayline_server
spica_shutter_server
spica_grating_server
spica_ipdet_server
spica_scidet_server
spica_power_server

## Slide 9
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
4. Templates
Templates are sequences of predefined elementary operations
A set of necessary scenarios for SPICA-VIS operation has been defined -> constraints on OS definition
Observation template
Calibration templates
Detector calibration
Spectral calibration
Kappa matrix calibration
Transfer function
Technical templates
Birefringence calibration
Reference pixels
SPICA-VIS / SPICA-FT cophasing
IMG/PUP command matrices
TTT command matrices

## Slide 10
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
5. OS : Principles and general organization
Communication with CHARA modules
(MF3)
GUI
(MF4)
Communication with SPICA/VIS servers
(MF2)
Scheduler
(MF1 / SF)
OS main functions
MF1: Scheduling operations for recording data (science)
MF2: Controlling all SPICA-VIS devices
MF3: Communicating with CHARA modules
MF4: Providing a Graphical User Interface (GUI)
OS secondary functions
SF1: Record of technical data
SF2: End of night operations
SF3: Diagnosis in case of a technical issue

## Slide 11
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
5. OS : Principles and general organization (continued)
Scheduler
Will be described in the next section
Graphical User Interface
  Observation general information
  Status of SPICA-VIS (and SPICA-FT)
  Scheduler status
  Errors, warnings, last relevant logs entries
  Real Time Displays (preliminary)
    IMG/PUP detector images
    Science detector images (full, interferometric or photometric channel)
    FT of interferometric channel : OPD, V², SNR, …
    Photometry monitoring (e.g. bargraphs)

## Slide 12
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
5. OS : Principles and general organization (continued)
Detectors software
Frame grabbing and copy in a shared memory
Several processes are connected to the shared memory (same or different pieces of software)
  Science detector -> 2 main processes
    Save frames (+ ancillary data) to FITS files
    Real time displays (as described before)
  IM/PUP detector -> 3 main processes
    Real time displays : images (or ROI) + photocenter computation during alignment procedure
    Servo of the 6 beams onto fiber : star position -> (PID + command matrix) -> TTT
    Save frames to FITS files (for technical purposes)
Logs
Need of timestamped logs for all SPICA-VIS servers and OS itself
Use of Linux multitail command to display last entries

## Slide 13
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
6. OS : Scheduler and sequences
Sequences
Corresponding to MF1 and SF functions and/or to the different templates
Scheduler principles
Scheduler is the core of OS : dedicated to automated
execution and control of the sequences flow
Each function (sequence) of OS is seen as a sequence
of several steps (macro-commands)
Each step is itself a list of elementary commands
Sequences and steps are described in script files
Command parser to decode and execute script files
Example of sequence : record of data science
Example (from VEGA record data script file) :
STDSP 1 CAMR_PORT1 PORT1 START_DET_PHOTON_RECORD
STDSP 1 CAMB_PORT1 PORT1 START_DET_PHOTON_RECORD
STDARSP 1 CAMR_PORT1 PORT1 GetNoFrame 1
STDARSP 1 CAMB_PORT1 PORT1 GetNoFrame 1
STDARSPX 1 CAMR_PORT1 PORT1 GET_RECORD_STATUS 1 1
STDARSPX 1 CAMB_PORT1 PORT1 GET_RECORD_STATUS 1 2
SOLVE_XnoVal 2 4
SLP 1000
PRINT RECORD_DATA_IN_PROGRESS
JMP -7
PRINT DATA_RECORD_DONE
SEND_MESSAGE_2_CHARA OBSERVE_FRINGE_DATA_DONE

## Slide 14
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
6. OS: Scheduler and sequences - an example : science data record sequence
Optimize flux injection and close TTT control loop
Align pupils and images
FRINGES record
NOFRINGES record
END
Start SPICA-FT
fringe tracking
RECORD_DONE
Find fringes Start tracking
START
Get record infos from CD
Check power and servers
OB infos
Config infos
Initialize
Wait for CHARA to be locked on star
STAR_ACQUIRED
DARK record
Initialize
SPICA-FT

## Slide 15
SPICA-VIS Design Review, July 23-24, 2020
SPICA-VIS Instrument Control Software and Observing Software
Frédéric Morand
7. SPICA-VIS messages
CHARA message | Direction
ASTROMOD_SET_STAR | From CD
ASTROMOD_SET_ACTIVE_SCOPES | From CD
ASTROMOD_SET_REFERENCE_CART | From CD
ASTROMOD_SET_OBS_WAVELENGTH | From CD
OBSERVE_PI_NAME | From CD
OBSERVE_PROGRAM_NAME | From CD
OBSERVE_OBJECT_TYPE | From CD
OBSERVE_STAR_ACQUIRED | From CD
OBSERVE_FRINGES_FOUND | From CD
OBSERVE_FRINGES_DATA_DONE | To CD
Use of CHARA messages to talk with other CHARA modules (mainly CD)
Need of creating several new messages specific to SPICA-VIS operations
Example of existing messages that OS would use